B2315P - Bluetooth Connection Protocol

Bluetooth parameters:

Service UUID: 0xFFF0Characteristic 1

UUID: 0xFFF1 Properties: Read | WriteCharacteristic 2

UUID: 0xFFF2 Properties: Notify

Protocol data structure:

Wearable devices are slaves, Bluetooth gateways, apps, etc. are masters
Wearable devices (slaves) sample data: heart rate, blood pressure, blood oxygen, body temperature, step counting, sleep. After establishing a connection with the host, the wearable device (slave) sends data to the host through Notify.

Note: B2315P defaults to blood oxygen once every 2 minutes, and the value is 0 when there is no blood oxygen.

Data structure:

| Byte Size | Format | Name | Unit | Description |
| --- | --- | --- | --- | --- |
| 1 | U8 | battery | / | 0–3 four levels, 0 is 0-25% (low battery), 1 is 25%-50%, 2 is 50%-75%, 3 is 75%-100 (full battery) |
| 1 | U8 | Hrs | Bpm | Heart rate |
| 1 | U8 | Sbp | mmHg | Blood pressure - Systolic blood pressure |
| 1 | U8 | Dbp | mmHg | Blood pressure - Diastolic blood pressure |
| 1 | U8 | Spo2 | % | Blood oxygen saturation |
| 2 | U16 | Wrist temperature | °C | Wrist temperature (Temp / 10 = actual wrist temperature) |
| 2 | U16 | Body temperature | °C | Body temperature (Temp / 10 = actual body temperature) |
| 4 | U8*3 | Steps | Step | Pedometer |
| 1 | U8 | Status | / | Bit0=Wearing status 1 for wearing, 0 for falling off, Bit1=Networked (not used, reserved for now), Bit2=SOS status, 1 for SOS status, 0 for normal status |

Example:

00426E4B610144016C00005301

00:0x00-->Hexadecimal to decimal-->0000--0-25% Low power

42:0x42-->Hexadecimal to Decimal-->66(Bpm, heart rate)

6E:0x6E-->Hexadecimal to Decimal-->110(mmHg, blood pressure - systolic)

4B:0x4B-->Hexadecimal to Decimal-->75(mmHg, blood pressure - diastolic)

61:0x61-->Hexadecimal to Decimal-->97(%, blood oxygen saturation)

0144:0x0144-->Hexadecimal to Decimal-->220-->220*0.1-->22.0（°C, wrist temperature）

016C:0x016C-->Hexadecimal to Decimal-->364-->364*0.1-->36.4 （°C, body temperature）

000053:0x000053-->Hexadecimal to Decimal-->83 83 steps

01:0x01-->Hexadecimal to Binary-->00000001   bit0=1  Wearing status, bit1=0 meaningless temporarily reserved bit2=0, 0 is normal state

Calibration time:

Service  UUID: 0xFFF0

Connect downlink command time synchronization: #time:yymmddhhmmss

Example: Year 2025, January 1, 1:01:01 #time:250101010101